## Slide 1
Third-party whole-slide image viewers do not produce the same image
Samuel Lam1, Qi Gong2, Wei-Chung Cheng2
1Department of Computer Science, University of Maryland, College Park; 2CDRH/OSEL/DIDSR
This study was supported by ORISE. The mention of commercial products herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services. We are grateful to Jonathan Boswell for helping in setting up this study.
ACKNOWLEDGEMENTS
METHODOLOGY
Sedeen:
Very close but not identical (µ=1.30, σ=1.35)
∆E pattern does not correlate with tissue structure
A few pixels have high ∆E (outliers)
ASAP and QuPath:
Different from reference (µ=18.79 and µ=18.69)
Different from each other (µ=2.19, σ=2.33) in microblock patterns
∆E correlates with tissue structure -- nuclei have higher ∆E (yellow); stroma lower (green)
ASAP and Sedeed
“Pixelized” patterns similar to lossy compression
FINDINGS
RESEARCH QUESTION
Do different WSI viewers generate identical images for the same WSI file?
Open the same WSI file with two different viewers
In the reference viewer, select the predefined region-of-interest (ROI)
Use AutoHotKey, a keyboard/mouse event generator, to select the same ROI automatically
Use Windows Snipping Tool to capture and save the screenshots [1]
Use Matlab to check the registration accuracy; if not registered correctly, go back to Step 3
Use Matlab to calculate the color difference (∆E) for each pixel
Use Matlab to report statistical and graphic results
RESULTS
Calvin Sun and Wei-Chung Cheng, How Much Can Bad Display Calibration Cost? 2017 FDA Summer Student Poster Day.
REFERENCES
ABSTRACT
Three freely available third-party whole-slide image viewers were compared with the factory viewer on the pixel level. Experiment results show that some viewers generated excessive errors.
BACKGROUND
A whole-slide imaging (WSI) system used in digital pathology consists of the scanner, viewer, and display components. So far, only two WSI devices have been cleared by the FDA. Recently, independent WSI viewers were submitted by third-party vendors to replace the original viewer component as alternatives. However, these viewers were not adequately tested because the vendors believe that the image is just digital data and will not be altered by their software.
TEST SUBJECTS
ASAP
Radboud U
NDP.view2
Hamamatsu
QuPath
Queen’s U
Sedeen
PathCore
Clinical Area: Medical Imaging
Experiments demonstrate why bench tests are required for third-party WSI viewers
Scanner
Viewer
Display
WSI
File
Input WSI file: “CMU-1.ndpi” generated by a Hamamatsu scanner
Four WSI viewers and their output images:
Sedeen
QuPath
NDP
ASAP
Four images: 6 pairs to compare
Heatmaps show dE for each pixel
Mean and std shown in title
Histogram shown in inset
Boxplot shown on right
WSI
Reference Viewer
3rd-party Viewer
AutoHotKey
Script
Screenshot
Registration
Color Difference
view
view*
image
image*
image
image*
∆E, heatmap,
histogram, boxplot
registration error
zoom/pan
CONCLUSIONS
For the same input file, four different WSI viewers generated four different images. It demonstrates the fallacy that any WSI viewer can reproduce digital images identically. The concept of interoperability between WSI components needs to be revisited to include image integrity on the pixel level. Before WSI file format is standardized, third-party viewer vendors should work with the scanner manufacturer instead of relying on untested free libraries. Adequate bench testing data are needed for 510(k) WSI viewer submissions.
Determine sources of ∆E (color profile? JPEG?)
Repeat experiments for cleared devices (Philips, Leica)
Develop bench test methods (software, hardware)
Determine acceptable criteria
Future Work