## Slide 1
Third-party whole-slide image viewers do not produce the same image
Samuel Lam1, Qi Gong2, Wei-Chung Cheng2
1University of Maryland, College Park; 2CDRH/OSEL/DIDSR
ABSTRACT
RESEARCH QUESTION
RESULTS
NDP.view2 and Sedeen generated similar images
ASAP and QuPath generated similar images
Large color differences between the two groups
Use of color profile is questionable
JPEG compression artifacts are prominent when comparing certain pairs of viewers.
Blank/void areas remains very similar between viewers.
FINDINGS
This study was supported by ORISE. The mention of commercial products herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services.
ACKNOWLEDGEMENTS
Determine testing data required to demonstrate performance equivalence between WSI viewer devices
REGULATORY RELEVANCE
BACKGROUND
METHODOLOGY
NDP.view2
ASAP
QuPath
Sedeen
In this study, a software program, the Image Viewer Integrity Evaluation System (IVIES), was developed to compare two image viewers on the pixel level and report any measured color differences. Four freely available image viewers were inspected with the test method. The results show that some of the viewers rendered the whole slide images differently compared with the factory image viewer.
A WSI system used in digital pathology consists of the scanner, image viewer, and display components. Recently, some independent image viewers were submitted by third-party companies to replace the factory image viewer component used in the predicate device. To demonstrate substantial equivalence between the former and the latter, sponsors were expected to provide bench testing data for review. However, the sponsors frequently did not test their image viewers adequately for image quality.
Do different WSI viewers generate identical images for the same WSI file?
Use different viewers to open the same WSI
Use a keyboard/mouse event generator to interact with the viewers to obtain the ROI
Capture the screenshot in Windows
Run registration; check registration accuracy
Calculate color difference ∆E for each pixel
Report statistics – mean, std, histogram, boxplot
NDP.View2 from Hamamatsu – reference
Sedeen from PathCore – DICOM WG26 chair
QuPath from Queen’s U
ASAP from Redboud U
SUBJECTS