## Slide 1
Third-party whole-slide image viewers do not produce the same image
Samuel Lam1, Qi Gong2, Wei-Chung Cheng3
1University of Maryland, College Park; 2CDRH/OSEL/DIDSR
ABSTRACT
RESEARCH QUESTION
NDP.view2 and Sedeen generated similar but not identical images
ASAP and QuPath generated similar images
Large color differences between the two groups
Use of color profile is questionable
JPEG compression artifacts are prominent when comparing certain pairs of viewers.
Blank/void areas remain very similar between viewers.
FINDINGS
This study was supported by ORISE. The mention of commercial products herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services. We are grateful to Jonathan Boswell for helping in setting up this study.
ACKNOWLEDGEMENTS
Experiments showing that testing data are required to demonstrate image integrity for 3rd-party WSI viewer software.
REGULATORY RELEVANCE
BACKGROUND
METHODOLOGY
In this study, three freely available 3rd-party image viewers were compared with the factory viewer. A software tool was developed to compare them on the pixel level. The results show that some of the viewers rendered the whole slide images very differently compared with the factory one.
A WSI system used in digital pathology consists of the scanner, image viewer, and display components. Recently, some independent image viewers were submitted by third-party companies to replace the factory image viewer component used in the predicate device. To demonstrate substantial equivalence between the former and the latter, sponsors were expected to provide bench testing data for review. However, the sponsors frequently did not test their image viewers adequately for image quality.
Do different WSI viewers generate identical images for the same WSI file?
AutoHotKey Script:
  Use different viewers to open the same WSI
  Use a keyboard/mouse event generator to interact with the viewers to obtain the ROI
  Capture the screenshot in Windows (this captures color information in the digital domain and does not involve the display) [1]
MATLAB:
  Run registration; check registration accuracy
  Calculate color difference ∆E for each pixel
  Report statistics (mean, standard deviation)
  Generate graphs for display
NDP.View2 from Hamamatsu (the reference)
Sedeen from PathCore
QuPath from Queen’s University Belfast
ASAP from Radboud University
TEST SUBJECTS
RESULTS
NDP.view2
ASAP
QuPath
Sedeen
Sun, C., Cheng, W-C., How Much Can Bad Display Cost? FDA Summer Student Poster Day 2018.
REFERENCES