## Slide 1
Third-party whole-slide image viewers do not produce the same image
Samuel Lam1, Qi Gong2, Wei-Chung Cheng3
1University of Maryland, College Park; 2CDRH/OSEL/DIDSR, 3CDRH/OSEL/DIDSR
In this study, a software program, the Image Viewer Integrity Evaluation System (IVIES), was developed to compare two image viewers on the pixel level and report any measured color differences. Four freely available image viewers were inspected with the test method. The results show that some of the viewers rendered the whole slide images differently compared with the factory image viewer.
ABSTRACT
Do different WSI viewers generate identical images for the same WSI file?
RESEARCH QUESTION
RESULTS
NDP.view2 and Sedeen are close in color
ASAP and QuPath are close in color
Color differences are quite large between other combinations (e.g. NDP.view2 and QuPath)
Compression artifacts are prominent when comparing certain pairs of viewers.
The color of the blank slide remains very similar between viewers.
FINDINGS
This study was supported by ORISE. The mention of commercial products herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services.
ACKNOWLEDGEMENTS
Determine testing data required to demonstrate performance equivalence between WSI viewer devices
REGULATORY RELEVANCE
A WSI system used in digital pathology consists of the scanner, image viewer, and display components. Recently, some independent image viewers were submitted by third-party companies to replace the factory image viewer component used in the predicate device. To demonstrate substantial equivalence between the former and the latter, sponsors were expected to provide bench testing data for review. However, the sponsors frequently did not test their image viewers adequately for image quality.
BACKGROUND
Use different viewers to open the same WSI
Use a keyboard/mouse event generator to interact with the viewers to obtain the ROI
Capture the screenshot in Windows
Run registration; check registration accuracy
Calculate color difference dE for each pixel
Report statistics
METHODOLOGY
NDP View from Hamamatsu – reference
Sedeen from PathCore – DICOM WG26 chair
QuPath from Queen’s U
ASAP from Redboud U
SUBJECTS