## Slide 1
Entanglement of color management components in a whole-slide imaging system
Wei-Chung Cheng, Chih-Lei Wu, Calvin Sun
CDRH/OSEL/DIDSR, U.S. Food and Drug Administration, Silver Spring, MD, USA 20993-0002
CORRELATION
METHODOLOGY
Within the computer environment, the pixel data generated by the review software is amended by three color management components -- display color profile, color manager, and graphics card software.
The display color profile is generated by the color
Three tools were used to track the pixel data: Snipping Tool (for RGBcaptured), Pixel Meter (for RGBdisplay), and spectroradiometer (for optical output).
Truth (best case): ΔE00=0; PCR=88.77% (<100%, when display calibrated perfectly)
Legacy WSI: ΔE00=22.35, midway between the best and worst cases; PCR=23.06%, lower scores from subject #2 and #3
Modern WSI: ΔE00=10.94, about 100% better than Legacy; PCR=70.35%, linearly correlated with Monochrome WSI and Truth (R2=0.9982)
Both PCR and ΔE00 can perfectly detect Legacy vs. Modern WSI devices
FINDINGS
Experiments were conducted to show how the software components affect the interoperability of the display in a whole-slide imaging (WSI) system.
ABSTRACT
SPIE Medical Imaging 2018,
Digital Pathology Conference,
February 13, 2017. Paper 10140-26
Fig. 1: The dataflow of a WSI system.
ACKNOWLEDGEMENT
This study was supported by Critical Path Initiative and ORISE. The mention of commercial products herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services.
CONCLUSION
Multispectral imaging technique enables the use of biological tissues as color targets for testing WSI devices.
The per-pixel color difference method quantitatively assesses color performance of WSI devices.
As a bench test, the per-pixel color difference method was justified by the psychophysical experiment data.
REFERENCES
[1] FDA, Guidance for WSI devices (2016).
[2] Y Yagi, Diagnostic Pathology 6 (suppl.1) (2011).
[3] WC Cheng et al., SPIE Medical Imaging (2013).
[4] P Shrestha et al., J Med Imaging 1.2 (2014).
[5] WC Revie et al., Anal Cell Pathol (2014).
[6] G. Sharma et al., Color Res. Appl. (2005).
BACKGROUND
A WSI system digitizes the glass slide and then reproduces the image for the human reader.
The scanner converts optical signals into digital signals. The display reverses the process.
Three parts are required for the scan process:
Scanner, scan software, and server-grade computer
Three parts are required for the review process:
Display, review software, and workstation-grade computer
The scan and review processes are connected by the WSI image file.
RESEARCH QUESTION
The scanner is highly complex and requires dedicated scan software and a powerful computer. These three parts are usually sold as a bundle.
The interoperability of review software and display is desired because:
hundreds of review workstations will be deployed in a hospital. The hardware cost of display and computer needs to be minimized, and
Third-party vendors offer more powerful review software.
Theoretically, the review software, computer, and display are interoperable if the interfaces are standardized.
For example, the standard interfaces of the WSI file format, operating system, display interface, and display color space can be DICOM Supplement 145, Windows, DisplayPort, and sRGB, respectively.
Research question: Does digital interoperability also guarantee optical interoperability?
Low-cost displays require color management to perform color correction, which is hidden from the software and user.
Our approach is to track the pixel data in the color management dataflow.